## Bids
Company | Amount | System | Meet Specs? | Delivery Days | Bid Form | Bond/Security | Non-Collusion | Spec Sheet | Notes
Ballistiglass LLC | 103400 | BallisDETECT |  | 21 | X | Cashier's Check | X | X
Cantronics | 129140 | Cantronics M3000SD3 |  | 29 | X | Cashier's Check | X | X
DNG | 102898.84 | CIC Technologies |  | 30 | X | Bond | X | X | Ballistiglass indicates DNG is accurate to .5F (Email 10:50am on 11.10.20)
Konica Minolta | 139458.18 | Mobotix |  | 10 | X | Bond | X | X
MCT Pro Tools | $287,810.00?? Illegible | ? |  | 20 | X | ? | X no county listed | ? | Non-Responsive
Point Security | 200499.73 | HI/PE Plus |  | by 12/30/2020 | X | Cashier's Check | X | X
Raptor | 263890 | Raptor Vision |  | 10 | X | Cashier's Check | X | X

## Spec Compare
Thermal Body Scanners |  |  | Ballisticglass | Cantronics | DNG | Konica | Point Security | Raptor
Specifications:
Sensor Type: Amorphous Silicon Microbolometer |  |  | Dual camera IR, Not specified thermal detector, either amphorus silicon or vanadium oxide. | Amorphous Silicon Microbolometer | Thermopile Array (Different) with 3 Sensors, Black body Control Required | Microbolometer | Multispectral Infrared: Walk Through Metal Detector with Temperature Scan (Different) | Amorphous Silicon Microbolometer
Thermal Resolution: 384 x 288 |  |  | 384 x 288 | 384x288 | 4x4 pixel or 100 PSI (Lower) | 336 x 252 (Lower) | Unknown | 480 x 360 (Better)
Spectral Response: 7-14μm |  |  | 700-1400nm (nanometers=nm), equal to 0.7-1.4μm (micrometer=μm); this ois a possible typo | 7-14μm | Unknown | 7.5-13.5μm (Lower) | 8-14μm (Lower) | Meets
Drift Compensation: Real-time continuous calibration (ext. TRS) |  |  | "Real-time" and missing continuous calibration and external TRS, which is a WHO standard | Real-time continuous calibration (ext. TRS) | Unknown | Unknown | Automatic | Real-time continuous calibration
Thermal Accuracy: ±0.18°F |  |  | ±0.10°F (Better) | ±0.18°F | ±0.15°F (Better) | ±0.20°F | ±0.54°F (Lower) | ±0.54°F (Lower)
Measurement Range: 86°F to 113°F |  |  | 70°F to 121°F (Better) | 86°F to 113°F | 75.2°F to 113°F (Better) | (-) 40°F to 320°F (Better) | 93°F to 108 °F (Lower) | Meets
Sensitivity: ≤50mK nominal (NETD) |  |  | Possible typo: <.50mk nominal | ≤50mK nominal | 0.15F = 83mK (Lower) | 50mk nominal | Unknown | Meets
Horizontal Field of View: 24.6° |  |  | 56° (Better) | 24.6° | 46° (Better) | 45°, 25°, or 17° (Better) | Unknown | 90° x 60° (Better)
Sensor Type: CMOS 1/3′′ Sensor |  |  | CMOS 1/3 | CMOS 1/3′′ Sensor | Unknown | CMOS 1/1.8" (Lower) | Unknown | Meets
Day Resolution: 2688 x 1520 (4MP) |  |  | "1920" ??? Missing Units | 2688 x 1520 (4MP) | 3840 x 2880 (11MP) (Better) | 3072 x 2048 (6MP) (Better) | Unknown | 1920 x 1080 (Lower)
Horizontal Field of View: Digitally synchronized to thermal image |  |  | No digital sycronization, "165 degrees" | Digitally synchronized to thermal image | Unknown | Unknown | Unknown | Meets
Frame Rate: Up to 50FPS |  |  | 60FPS (Better) | Up to 50FPS | 60FPS (Better) | Unknown | Up to 50FPS | Meets
Contrast: Automatic |  |  | Automatic | 680 x 480 on screen/processing, contrast and brightness adjustments are automatic | Automatic | Unknown | Unknown | Meets
Brightness: Automatic |  |  | Automatic | 680 x 480 on screen/processing, contrast and brightness adjustments are automatic | Automatic | Unknown | Unknown | Meets
Isotherm Color Display: Pseudocolour specifically for fever screening |  |  | False Color | User definable isotherm color display for fever screening and alarm trigger thresholds | Unknown | False colors and black or white | Unknown | Meets
Temperature Trigger: User-defined temperature trigger threshold |  |  | User defined | User definable isotherm color display for fever screening and alarm trigger thresholds | Unknown | User definable 20 temperature triggers | Prebuilt Programs | Meets
Febrile Tracking: Yes (Synchronized dual-sensor tracking) |  |  | Yes, both cameras | Digitally synchronized to thermal image field of view | Unknown | Unknown | Unknown | Meets
Zoning: Yes |  |  | Yes | Yes | Unknown | Yes | Unknown | Meets
Masking: Yes |  |  | Yes | Yes | Unknown | Yes | Unknown | Meets
Alarm: Audiovisual with intelligent febrile tracking |  |  | Sound, visual, relay | Ignore mode or save alarm mode, audiovisual with intelligent febrile tracking, sound/visual/network alarm, user selectable email notification of alarm triggered event (Better) | Correlated video cam output with thermal scan data, target temperature is determined (Different) | Unknown | Walk Through Metal Detector with Temperature Scan, 50 built-in programs, programmable for 93°F to 108 °F (Different) | Meets: Audio and Visual Alarms
Storage: Micro SD Card (Up to 256 GB) or/and NVR |  |  | Included laptop with solid state hard drive; NVR (Network video recording) not specified so only at source recording, storage capacity not specified. | Micro SD Card (Up to 256 GB) or/and NVR | Unknown | Micro SD Card (4GB) | Unknown | SD and CPU Storage
Recording Output: 1080p |  |  | 1080p | 1080p | "YES" | 6MP (Better) | Unknown | Meets
USB Connector: USB 3.0 x 1, USB 2.0 x 2 |  |  | Only "USB 3.0", asuming one; two USB 2.0 not indicated. This specification indicates three ports. | USB 3.0 x 1, USB 2.0 x 2 | Unknown | Unknown | Unknown | HDMI
Input Voltage: 110 – 230V AC |  |  | 110 – 230V AC | 110 – 230V AC | Unknown | Unknown | 100 – 240V AC (Higher) | Meets
Operating Temperature: 60°F to 91°F |  |  | 50°F to 94°F (Better) | 50°F to 95°F (Better) | Unknown | Unknown | 50°F to 86°F (Lower) | Meets
Storage Temperature: -40°F to 185°F |  |  | Not specified-missing | -40°F to 185°F | Unknown | -40°F to 140°F (Lower) | (-)35°F to 158°F (Lower) | 32°F to 104°F (Better)
Installation Options: Wall Mounted, Ceiling Mounted or Mobile Pedestal |  |  | Camera Tripod | Wall Mounted, Ceiling Mounted or Mobile Pedestal | Unknown | Wall, pole or ceiling mounted | No, walk through detector | Wall, tripod, and pole mounted
Extra features (Unsolicited) |  |  | N/A | Added features of (multi) network support protocols (Better) | "YES" Tablet Support Wi-Fi | Integrated Microphone & Speaker, PoE Thermal Camera | N/A | Facial recognition of 99.8% accuracy within 8.2 FT, 20 plus people at one time.
| White | No Clarification or Data Sheet
| Yellow | Data Not Provided | Security Glass Company: CE Certification is European Certification, Camera offered is China sourced. | Thermal Scan Company (Most Responsive) | Animal Tracking Outdoors | Security Camera | Indoor Metal Detector | Large Volume Visitor Scanning
| Green | Good as Submitted
| D. Green | Better than Required
| Red | Not Adequate to Spec
| Blue | Diversion from Specification